## Read me
Purpose | Portable review of v2 claims and measurement-chain relationships (prototype: K143272).
Edit only | Yellow columns (review decision, reviewed value/category, reviewer, date, notes). Import rejects any change to IDs, values or fingerprints.
Decisions | pending · verified · edited and verified · rejected · needs clarification
Rules | Edit-and-verify needs a corrected value or category. Needs-clarification needs a note. Verifying an inferred item needs a note. A 'link not stated' relationship cannot be verified as a link.
Table 8.3 | Verified and edited-and-verified, non-stale claims connected through verified stated relationships only. Clearance and family counts are separate.
Staleness | When evidence changes after review, the evidence fingerprint changes and the review stops counting until it is re-confirmed.
Import | node scripts/v2/import-review.mjs <this file>.xlsx [--check-only]
Authority | The original FDA PDF is the ultimate regulatory source. This workbook is an unofficial research derivative.

## Claims review
Claim ID | 510(k) | Chain(s) | Field | Value state | Polarity | Proposed normalized value | Table 8.3 category | Source wording | Interpretation | Evidence scope | Page | Primary exact quote | Evidence fingerprint | Review decision | Reviewed value | Reviewed category | Reviewer | Reviewer date | Reviewer notes | Reviewed evidence fingerprint | Evidence page
CLM-5f420010811a | K143272 |  | indications_for_use | stated | affirmed | The ApneaLink™ Air device is indicated for use by Health Care Professionals (HCP), where it may aid in the diagnosis of sleep disordered breathing for adult patients. ApneaLink Air records the following data: patient respiratory nasal airflow, snoring, blood oxygen saturation, pulse, respiratory effort and body position during sleep. The device uses these recordings to produce a report for the HCP that may aid in the diagnosis of sleep disordered breathing or for further clinical investigation. The device is intended for home and hospital use under the direction of a HCP. |  | The ApneaLink™ Air device is indicated for use by Health Care Professionals (HCP), where it may aid in the diagnosis of sleep disordered breathing for adult patients. ApneaLink Air records the following data: patient respiratory nasal airflow, snoring, blood oxygen saturation, pulse, respiratory effort and body position during sleep. The device uses these recordings to produce a report for the HCP that may aid in the diagnosis of sleep disordered breathing or for further clinical investigation. The device is intended for home and hospital use under the direction of a HCP. | explicit | subject | 5 | The ApneaLink™ Air device is indicated for use by Health Care Professionals (HCP), where it may aid in the diagnosis of sleep disordered breathing for adult patients. ApneaLink Air records the following data: patient respiratory nasal airflow, snoring, blood oxygen saturation, pulse, respiratory effort and body position during sleep. The device uses these recordings to produce a report for the HCP that may aid in the diagnosis of sleep disordered breathing or for further clinical investigation. The device is intended for home and hospital use under the direction of a HCP. | 40ccf26ae4699748 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-5f420010811a
CLM-69b71b898d7c | K143272 |  | intended_use_role | stated | affirmed | aid to diagnosis |  | aid in the diagnosis | normalized | subject | 5 | where it may aid in the diagnosis of sleep disordered breathing for adult patients | face8ad7bee763ec | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-69b71b898d7c
CLM-39b492a99694 | K143272 |  | intended_use_role | stated | affirmed | evaluation |  | aid in the diagnosis | contradicted | subject | 5 | where it may aid in the diagnosis of sleep disordered breathing for adult patients | 6ec62089ea4a7fc1 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-39b492a99694
CLM-006338992a5c | K143272 | CHN-K143272-AIRFLOW | sensor_technology | stated | affirmed | Solid-state pressure sensor | Nasal pressure transducer | Solid-state pressure sensor | explicit | subject | 9 | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels | 9e0335036458925e | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-006338992a5c
CLM-9feb9ae5aa69 | K143272 | CHN-K143272-AIRFLOW | sensor_location | stated | affirmed | Nasal cannula at patient's nose | Nares / nasal cannula | Nasal cannula at patient's nose | explicit | subject | 5 | The disposable plastic nasal cannula is connected to the ApneaLink Air recorder and fixed at the patient’s nose. | 4d2f79f73614aacc | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-9feb9ae5aa69
CLM-fb528aac2fe3 | K143272 | CHN-K143272-AIRFLOW | raw_acquired_signal | stated | affirmed | Nasal pressure | Nasal pressure signal (waveform not specified) | Nasal pressure | explicit | subject | 8 | flow channel for measurement of patient respiratory nasal pressure. | 5e8256c761258529 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-fb528aac2fe3
CLM-b7527e172867 | K143272 | CHN-K143272-AIRFLOW | physiological_parameter | stated | affirmed | Respiratory nasal airflow | Airflow / respiration | Respiratory nasal airflow | explicit | subject | 5 | patient respiratory nasal airflow | 8834f0efa8d5bb5a | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-b7527e172867
CLM-25cea7da6112 | K143272 |  | sensor_technology | stated | affirmed | Respiratory effort sensor (RIP technology) |  | Respiratory effort sensor (RIP technology) | contradicted | subject | 9 | Respiratory Effort Sensors (RIP technology) | d9d6ac50f5d418c9 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-25cea7da6112
CLM-64a1213f973f | K143272 | CHN-K143272-EFFORT | sensor_location | stated | affirmed | Patient's chest parallel to spine | Chest | Patient's chest parallel to spine | explicit | subject | 5 | The ApneaLink Air recorder and the respiratory effort sensor must be fastened with the re-usable belt on the patient’s chest parallel to the spine. | 176eea0c6a004a64 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-64a1213f973f
CLM-e36feb851fad | K143272 |  | raw_acquired_signal | stated | affirmed | Respiratory effort |  | Respiratory effort | contradicted | subject | 8 | effort channel to measure the respiratory effort using a pneumatic principle. | 2033c3b09e5d11b2 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-e36feb851fad
CLM-cb110eefecb7 | K143272 | CHN-K143272-EFFORT | physiological_parameter | stated | affirmed | Respiratory effort | Respiratory effort | Respiratory effort | explicit | subject | 8 | effort channel to measure the respiratory effort using a pneumatic principle. | 68070619bbb2b72e | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-cb110eefecb7
CLM-0924252480a1 | K143272 | CHN-K143272-EFFORT | direct_vs_derived | stated | affirmed | measured | Directly sensed | measured | normalized | subject | 8 | effort channel to measure the respiratory effort using a pneumatic principle. | 7e3435c95dcc43f3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-0924252480a1
CLM-73e2718fb82d | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | sensor_technology | stated | affirmed | Oximetry sensor | Oximetry sensor, technology unspecified | Oximetry sensor | explicit | subject | 9 | Utilizes oximetry sensor | 1382a2de83c9b84c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-73e2718fb82d
CLM-bf8541d0dc67 | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | sensor_location | stated | affirmed | Patient's finger | Finger | Patient's finger | explicit | subject | 5 | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger. | 065090f63ea2fb65 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-bf8541d0dc67
CLM-b02ad0500a1b | K143272 |  | raw_acquired_signal | stated | affirmed | Oxygen saturation and pulse rate |  | Oxygen saturation and pulse rate | contradicted | subject | 8 | oximeter input to measure degree of oxygen saturation of the blood and pulse rate. | 2078d3a02b50b73d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-b02ad0500a1b
CLM-f706d25f6cda | K143272 |  | physiological_parameter | stated | affirmed | Blood oxygen saturation and pulse |  | Blood oxygen saturation and pulse | normalized | subject | 8 | oximeter input to measure degree of oxygen saturation of the blood and pulse rate. | 12d4c38d2f7f6692 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-f706d25f6cda
CLM-eca11ab54826 | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | direct_vs_derived | stated | affirmed | measured | Directly sensed | measured | normalized | subject | 8 | oximeter input to measure degree of oxygen saturation of the blood and pulse rate. | faa60e00aa77cfba | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-eca11ab54826
CLM-5db33eba2fd1 | K143272 | CHN-K143272-POSITION | sensor_technology | stated | affirmed | Actigraphy sensor | Actigraphy sensor, technology unspecified | Actigraphy sensor | explicit | subject | 9 | Utilizes solid-state actigraphy sensor for position | 24d4ab08e2dfb135 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-5db33eba2fd1
CLM-32b036c02477 | K143272 | CHN-K143272-POSITION | sensor_location | stated | affirmed | Inside recorder worn on chest | Chest (within recorder; inferred) | Inside recorder worn on chest | inferred | comment | 8 | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated. | f2d814023f943ffc | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-32b036c02477
CLM-a5dcc1a3a694 | K143272 |  | raw_acquired_signal | stated | affirmed | Body position |  | Body position | contradicted | comment | 8 | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated. | 0c7b079a3ead209e | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-a5dcc1a3a694
CLM-78792476b8f4 | K143272 | CHN-K143272-POSITION | physiological_parameter | stated | affirmed | Body position | Body position | Body position | explicit | subject | 8 | body position channel to determine body position of the patient during sleep | 0e55fff859b56ee6 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-78792476b8f4
CLM-964ea01d51e5 | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | AHI | AHI | AHI | explicit | shared | 11 | AHI | ccb081a2fa87bc03 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-964ea01d51e5
CLM-2102f8b48d70 | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | RI | RI | RI | explicit | shared | 11 | RI | 6895e007594ced13 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-2102f8b48d70
CLM-f2dd379ff37c | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Apnea Index | Apnea index | Apnea Index | explicit | shared | 11 | Apnea Index (unclassified, central, mixed, obstructive) | ff27da481048e9b8 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-f2dd379ff37c
CLM-aeebd6206c8c | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Hypopnea Index | Hypopnea index | Hypopnea Index | explicit | shared | 11 | Hypopnea Index | 3376bf233e2ad888 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-aeebd6206c8c
CLM-7cdb877b6e99 | K143272 | CHN-K143272-AIRFLOW, CHN-K143272-REPORT, CHN-K143272-SNORING | derived_reported_feature | stated | affirmed | Number of flow limited breaths without snoring | Flow limitation | Number of flow limited breaths without snoring | explicit | shared | 11 | Number of flow limited breaths without snoring | 36057b9d0a46952b | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-7cdb877b6e99
CLM-9bfd0bef8160 | K143272 | CHN-K143272-AIRFLOW, CHN-K143272-REPORT, CHN-K143272-SNORING | derived_reported_feature | stated | affirmed | Number of flow limited breaths with snoring | Flow limitation with snoring | Number of flow limited breaths with snoring | explicit | shared | 11 | Number of flow limited breaths with snoring | 60169e2d01ae58d8 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-9bfd0bef8160
CLM-ac3ad056bb9e | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Cheyne-Stokes Breathing | Cheyne-Stokes breathing | Cheyne-Stokes Breathing | explicit | shared | 11 | Cheyne-Stokes Breathing | b1cd574498d7e9ff | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-ac3ad056bb9e
CLM-1cdb0340ac87 | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | ODI | ODI | ODI | explicit | shared | 11 | ODI | d63acb928ed1e5c1 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-1cdb0340ac87
CLM-e7ba308a04c1 | K143272 | CHN-K143272-REPORT, CHN-K143272-SPO2 | derived_reported_feature | stated | affirmed | Average saturation | SpO2 average | Average saturation | explicit | shared | 11 | Average saturation | c2780759e74b2020 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-e7ba308a04c1
CLM-5de192f8ed59 | K143272 | CHN-K143272-REPORT, CHN-K143272-SPO2 | derived_reported_feature | stated | affirmed | Minimum saturation | SpO2 minimum | Minimum saturation | explicit | shared | 11 | Minimum saturation | a8d3163f89cc0370 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-5de192f8ed59
CLM-83c2442cfd22 | K143272 | CHN-K143272-REPORT, CHN-K143272-SPO2 | derived_reported_feature | stated | affirmed | Maximum saturation | SpO2 maximum | Maximum saturation | explicit | shared | 11 | Maximum saturation | bc02d69373f1c409 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-83c2442cfd22
CLM-e1cd3ac8058c | K143272 | CHN-K143272-REPORT, CHN-K143272-SPO2 | derived_reported_feature | stated | affirmed | Basal saturation | SpO2 basal | Basal saturation | explicit | shared | 11 | Basal saturation | 158e2691186f910c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-e1cd3ac8058c
CLM-a9ce2d729da8 | K143272 | CHN-K143272-PULSE, CHN-K143272-REPORT | derived_reported_feature | stated | affirmed | Minimum Pulse Rate | Pulse rate minimum | Minimum Pulse Rate | explicit | shared | 11 | Minimum Pulse Rate | d4f35ac78d1b8303 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-a9ce2d729da8
CLM-5a17f0947641 | K143272 | CHN-K143272-PULSE, CHN-K143272-REPORT | derived_reported_feature | stated | affirmed | Maximum Pulse Rate | Pulse rate maximum | Maximum Pulse Rate | explicit | shared | 11 | Maximum Pulse Rate | 78fe419555df3440 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-5a17f0947641
CLM-f246d466ecda | K143272 | CHN-K143272-PULSE, CHN-K143272-REPORT | derived_reported_feature | stated | affirmed | Average Pulse Rate | Pulse rate average | Average Pulse Rate | explicit | shared | 11 | Average Pulse Rate | 4f8815325072d932 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-f246d466ecda
CLM-030e4da72b91 | K143272 |  | derived_reported_feature | stated | affirmed | Body position data |  | Body position data | contradicted | subject | 5 | Data is neither part of the report generated by ApneaLink Software nor does it influence or alter any index of the report. | 3ac7fd6f1d14becb | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-030e4da72b91
CLM-77cb9c564d4e | K143272 | CHN-K143272-AIRFLOW | sensor_technology | stated | affirmed | Nasal cannula with plastic tubing (pressure-sensing patient interface) |  | Plastic tubing and cannula for pressure sensing | explicit | subject | 8 | Plastic tubing and cannula for pressure sensing | c0767787b900e0c1 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-77cb9c564d4e
CLM-a7cf317deee5 | K143272 | CHN-K143272-AIRFLOW | sensor_location | Not stated | affirmed | Not stated |  |  | not stated | subject | 5 | The disposable plastic nasal cannula is connected to the ApneaLink Air recorder and fixed at the patient’s nose. | 139edea41e10bd7b | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-a7cf317deee5
CLM-28294e7255aa | K143272 | CHN-K143272-AIRFLOW | direct_vs_derived | stated | affirmed | Derived: nasal pressure values are transformed (linearized) into flow values in the ApneaLink Air recorder | Derived | Transformation of pressure values in flow values | normalized | subject | 10 | ApneaLink Air recorder | bb146f295f5e7625 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-28294e7255aa
CLM-0be59426dade | K143272 | CHN-K143272-EFFORT | sensor_technology | stated | affirmed | Pneumatic respiratory effort sensor (solid-state pressure sensor) | Pneumatic effort sensor | effort channel ... using a pneumatic principle | normalized | subject | 8 | effort channel to measure the respiratory effort using a pneumatic principle. | 879f0d2099ab9123 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-0be59426dade
CLM-5f8f25c65423 | K143272 | CHN-K143272-EFFORT | raw_acquired_signal | stated | affirmed | Pressure changes from the pneumatic effort sensor, converted to electrical signal levels | Pneumatic pressure signal (waveform not specified) | converts pressure changes to electrical signal levels | normalized | subject | 9 | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels | 0a512983b4af3d13 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-5f8f25c65423
CLM-6054f03c13fa | K143272 | CHN-K143272-EFFORT | calibration | Not stated | affirmed | Not stated |  |  | not stated |  |  |  | e92560d6f86bcdef | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-6054f03c13fa
CLM-54f39c86b89c | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | sensor_technology | stated | affirmed | Pulse oximeter module: Nonin Xpod 3012 LP (low-power Xpod) |  | Nonin Xpod 3012 LP | explicit | subject | 13 | Nonin Xpod 3012 LP | 06968a492aa2a307 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-54f39c86b89c
CLM-e6276e6057cf | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | sensor_location | stated | affirmed | Xpod module fixed at the belt via clip |  | fixed at the belt via clip | explicit | subject | 8 | The new low power Xpod is fixed at the belt via clip. | ab6f41e0963ea7e0 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-e6276e6057cf
CLM-6d87e8e5044c | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | raw_acquired_signal | Not stated | affirmed | Not stated | Not specified |  | not stated | subject | 8 | oximeter input to measure degree of oxygen saturation of the blood and pulse rate. | 4f2bec0ff7392c8c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-6d87e8e5044c
CLM-067ec2b7756b | K143272 | CHN-K143272-SPO2 | physiological_parameter | stated | affirmed | Blood oxygen saturation (SpO2) | Oxygen saturation | oxygen saturation of the blood | explicit | subject | 8 | oximeter input to measure degree of oxygen saturation of the blood and pulse rate. | fe3b34daf10f169a | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-067ec2b7756b
CLM-92a1c7763280 | K143272 | CHN-K143272-PULSE | physiological_parameter | stated | affirmed | Pulse rate | Cardiac activity / heart rate | pulse rate | explicit | subject | 8 | oximeter input to measure degree of oxygen saturation of the blood and pulse rate. | 8918504c70c5910d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-92a1c7763280
CLM-a8c39fbe3950 | K143272 | CHN-K143272-PULSE, CHN-K143272-SPO2 | validation_comparator | stated | affirmed | Low-power Xpod stated to have the same accuracy and performance as the previous Xpod; test method and comparator not stated |  | Same accuracy and performance as the existing Xpod | normalized | subject | 5 | Same accuracy and performance as the existing Xpod, but considerably lower power draw. | 472f2799c97c5803 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-a8c39fbe3950
CLM-78429ae1bb69 | K143272 | CHN-K143272-POSITION | raw_acquired_signal | stated | affirmed | Device tilt/rotation angle | Orientation / angle signal | the angle the device is tilted or rotated | normalized | comment | 8 | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated. | b26849c016e8b4b3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-78429ae1bb69
CLM-a6f595eac9ad | K143272 | CHN-K143272-POSITION | direct_vs_derived | stated | affirmed | Derived from device orientation (tilt/rotation angle) | Derived | according to the angle the device is tilted or rotated | normalized | comment | 8 | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated. | 8fdbeb0e09712743 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-a6f595eac9ad
CLM-e0d8610a08dd | K143272 | CHN-K143272-POSITION | diagnostic_output | stated | negated | Not part of the ApneaLink Software report (explicit exclusion) | Excluded from report | neither part of the report generated by ApneaLink Software | explicit | subject | 5 | Data is neither part of the report generated by ApneaLink Software nor does it influence or alter any index of the report. | 95d7236892dd55ff | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-e0d8610a08dd
CLM-7ec48c11db89 | K143272 | CHN-K143272-POSITION | signal_access | stated | affirmed | Retained and accessible: stored as an EDF+ file on the recorder; readable via USB with an EDF viewer | Yes | stored as an EDF+ file on the recorder | explicit | subject | 5 | The body position data is stored as an EDF+ file on the recorder. It is accessible on a computer via USB connection using an EDF viewer. | 8e0d2a7a34d5c123 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-7ec48c11db89
CLM-55d0a82ea239 | K143272 | CHN-K143272-POSITION | validation_comparator | stated | affirmed | Side-by-side testing against Nox T3 using a protractor to show defined positions around the body axis |  | side by side testing ... using a protractor | explicit | subject | 15 | Additional side by side testing has been performed by using a protractor to show the defined position around the body axis for ApneaLink Air and Nox T3. | 34ec43dc7c2f170c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-55d0a82ea239
CLM-fae298439e0a | K143272 | CHN-K143272-SNORING | physiological_parameter | stated | affirmed | Snoring | Snoring / respiratory sounds | snoring | explicit | subject | 5 | patient respiratory nasal airflow, snoring, blood oxygen saturation | 0a30b72cd1312299 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-fae298439e0a
CLM-f923ac3ff49e | K143272 | CHN-K143272-SNORING | sensor_technology | Not stated | affirmed | Not stated | Sensor not stated |  | not stated |  |  |  | 241974da55b1c48f | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-f923ac3ff49e
CLM-3b26fd752298 | K143272 | CHN-K143272-SNORING | sensor_location | Not stated | affirmed | Not stated | Not specified |  | not stated |  |  |  | 8dc1e6e41bc4ab86 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-3b26fd752298
CLM-e00e5a54595a | K143272 | CHN-K143272-SNORING | raw_acquired_signal | Not stated | affirmed | Not stated | Not specified |  | not stated |  |  |  | 8dc1e6e41bc4ab86 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-e00e5a54595a
CLM-512244c68e62 | K143272 | CHN-K143272-SNORING | direct_vs_derived | Not stated | affirmed | Not stated | Not stated |  | not stated |  |  |  | f4bf8f55eb949fe5 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-512244c68e62
CLM-f15b1f62480b | K143272 | CHN-K143272-AIRFLOW, CHN-K143272-EFFORT, CHN-K143272-PULSE, CHN-K143272-SPO2 | signal_access | stated | affirmed | Unclear: recorded data are stored as EDF+ on the SD card and readable over USB as mass storage, but the channels and waveform content are not itemized | Unclear | Recorded data is stored in European Data Format (EDF+) | normalized | subject | 10 | Recorded data is stored in European Data Format (EDF+) file on SD card. When device is connected to PC via USB the device provides access to its internal memory as mass storage memory including the recorded EDF+ files. | 563c1b34ab70918c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-f15b1f62480b
CLM-fcdeccd52d61 | K143272 | CHN-K143272-AIRFLOW, CHN-K143272-EFFORT, CHN-K143272-POSITION, CHN-K143272-PULSE, CHN-K143272-SPO2 | sampling_resolution | Not stated | affirmed | Not stated |  |  | not stated |  |  |  | e92560d6f86bcdef | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-fcdeccd52d61
CLM-a0f1495fa55b | K143272 | CHN-K143272-AIRFLOW, CHN-K143272-EFFORT, CHN-K143272-PULSE, CHN-K143272-SPO2 | continuity | Not stated | affirmed | Not stated |  |  | not stated |  |  |  | e92560d6f86bcdef | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-a0f1495fa55b
CLM-6036fea343be | K143272 | CHN-K143272-AIRFLOW, CHN-K143272-EFFORT, CHN-K143272-PULSE, CHN-K143272-SPO2 | synchronization | Not stated | affirmed | Not stated |  |  | not stated |  |  |  | e92560d6f86bcdef | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-6036fea343be
CLM-0e8dbbca0f00 | K143272 | CHN-K143272-REPORT | validation_comparator | stated | affirmed | Event detection and indices not re-tested here: cleared under K083575 (ApneaLink Plus); ApneaLink Pro side-by-side results relied on; comparator not described in this summary |  | Detection of respiratory events and reported indices was cleared under K083575 | normalized | subject | 15 | Detection of respiratory events and reported indices was cleared under K083575 (ApneaLink Plus) and remained unchanged in ApneaLink Pro (K131932) and ApneaLink Air. | c8865d1a10064b73 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-0e8dbbca0f00
CLM-8ac34c8c8f77 | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Report for the HCP generated by the ApneaLink Software | HCP report | produce a report for the HCP | explicit | subject | 5 | The device uses these recordings to produce a report for the HCP | 0d29c97d9b577ef1 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-8ac34c8c8f77
CLM-78f52f4babfa | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Apnea Index, unclassified | Apnea index (unclassified) | unclassified | normalized | shared | 11 | Apnea Index (unclassified, central, mixed, obstructive) | d3e4c9f0bf2b564d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-78f52f4babfa
CLM-b0c786c21ea2 | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Apnea Index, central | Apnea index (central) | central | normalized | shared | 11 | Apnea Index (unclassified, central, mixed, obstructive) | bf51677510f1bb89 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-b0c786c21ea2
CLM-8cee7033615b | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Apnea Index, mixed | Apnea index (mixed) | mixed | normalized | shared | 11 | Apnea Index (unclassified, central, mixed, obstructive) | 3caa3013b7dfabc8 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-8cee7033615b
CLM-bcc4fe94e099 | K143272 | CHN-K143272-REPORT | diagnostic_output | stated | affirmed | Apnea Index, obstructive | Apnea index (obstructive) | obstructive | normalized | shared | 11 | Apnea Index (unclassified, central, mixed, obstructive) | d0e6be3a14372d1d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=CLM-bcc4fe94e099

## Relationships review
Relationship ID | 510(k) | Chain | From claim | From value | Relation | To claim | To value | Basis | Note | Evidence fingerprint | Review decision | Reviewed value | Reviewed category | Reviewer | Reviewer date | Reviewer notes | Reviewed evidence fingerprint | Evidence page
REL-99a63111d1a2 | K143272 | CHN-K143272-AIRFLOW | CLM-77cb9c564d4e | Nasal cannula with plastic tubing (pressure-sensing patient interface) | component_of | CLM-006338992a5c | Solid-state pressure sensor | explicit |  | 510a9313f8741ebf | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-99a63111d1a2
REL-6669ab69f70e | K143272 | CHN-K143272-AIRFLOW | CLM-77cb9c564d4e | Nasal cannula with plastic tubing (pressure-sensing patient interface) | located_at | CLM-9feb9ae5aa69 | Nasal cannula at patient's nose | explicit |  | 06cd97bdae50c554 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-6669ab69f70e
REL-685ed881f59a | K143272 | CHN-K143272-AIRFLOW | CLM-006338992a5c | Solid-state pressure sensor | located_at | CLM-a7cf317deee5 | Not stated | not-stated | Transducer location not stated. | 330685b63b76a067 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-685ed881f59a
REL-4612222a5926 | K143272 | CHN-K143272-AIRFLOW | CLM-006338992a5c | Solid-state pressure sensor | acquires | CLM-fb528aac2fe3 | Nasal pressure | explicit |  | 2c23b21988821766 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-4612222a5926
REL-7fa8540c4751 | K143272 | CHN-K143272-AIRFLOW | CLM-fb528aac2fe3 | Nasal pressure | represents | CLM-b7527e172867 | Respiratory nasal airflow | explicit | Pressure values are transformed into flow values. | 7bf7afe821fe88ad | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-7fa8540c4751
REL-9734ff192882 | K143272 | CHN-K143272-AIRFLOW | CLM-b7527e172867 | Respiratory nasal airflow | derivation | CLM-28294e7255aa | Derived: nasal pressure values are transformed (linearized) into flow values in the ApneaLink Air recorder | explicit |  | 414972fe16e1d89f | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-9734ff192882
REL-79d1e425d13a | K143272 | CHN-K143272-AIRFLOW | CLM-7cdb877b6e99 | Number of flow limited breaths without snoring | derived_feature_of | CLM-b7527e172867 | Respiratory nasal airflow | output-name | The output name refers to 'flow'. The input signal is not otherwise stated. | b1985baa1b55c68e | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-79d1e425d13a
REL-e8498d05cc18 | K143272 | CHN-K143272-AIRFLOW | CLM-9bfd0bef8160 | Number of flow limited breaths with snoring | derived_feature_of | CLM-b7527e172867 | Respiratory nasal airflow | output-name | The output name refers to 'flow'. The input signal is not otherwise stated. | 023faecb1769e75b | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-e8498d05cc18
REL-2fe3896d9f1a | K143272 | CHN-K143272-AIRFLOW | CLM-964ea01d51e5 | AHI | reported_as | CLM-b7527e172867 | Respiratory nasal airflow | not-stated | Which signals this index uses is not stated in the summary. | 0a4d1a667c728ddd | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-2fe3896d9f1a
REL-f967d8f13686 | K143272 | CHN-K143272-AIRFLOW | CLM-2102f8b48d70 | RI | reported_as | CLM-b7527e172867 | Respiratory nasal airflow | not-stated | Which signals this index uses is not stated in the summary. | 0a4d1a667c728ddd | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f967d8f13686
REL-da1a5564e93c | K143272 | CHN-K143272-AIRFLOW | CLM-f2dd379ff37c | Apnea Index | reported_as | CLM-b7527e172867 | Respiratory nasal airflow | not-stated | Which signals this index uses is not stated in the summary. | 0a4d1a667c728ddd | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-da1a5564e93c
REL-597c311bfac2 | K143272 | CHN-K143272-AIRFLOW | CLM-aeebd6206c8c | Hypopnea Index | reported_as | CLM-b7527e172867 | Respiratory nasal airflow | not-stated | Which signals this index uses is not stated in the summary. | 0a4d1a667c728ddd | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-597c311bfac2
REL-aa411550ac0a | K143272 | CHN-K143272-AIRFLOW | CLM-ac3ad056bb9e | Cheyne-Stokes Breathing | reported_as | CLM-b7527e172867 | Respiratory nasal airflow | not-stated | Which signals this index uses is not stated in the summary. | 0a4d1a667c728ddd | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-aa411550ac0a
REL-e8750cd5dc13 | K143272 | CHN-K143272-EFFORT | CLM-0be59426dade | Pneumatic respiratory effort sensor (solid-state pressure sensor) | located_at | CLM-64a1213f973f | Patient's chest parallel to spine | explicit |  | 9efda5c8e5253400 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-e8750cd5dc13
REL-70a59a69c068 | K143272 | CHN-K143272-EFFORT | CLM-0be59426dade | Pneumatic respiratory effort sensor (solid-state pressure sensor) | acquires | CLM-5f8f25c65423 | Pressure changes from the pneumatic effort sensor, converted to electrical signal levels | explicit | Shared cell: 'for nasal flow and effort'. | 277c024f772e76a4 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-70a59a69c068
REL-b58c7986c74e | K143272 | CHN-K143272-EFFORT | CLM-5f8f25c65423 | Pressure changes from the pneumatic effort sensor, converted to electrical signal levels | represents | CLM-cb110eefecb7 | Respiratory effort | explicit |  | e64f4a4f43d04bcf | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-b58c7986c74e
REL-be77cedaa7a8 | K143272 | CHN-K143272-EFFORT | CLM-cb110eefecb7 | Respiratory effort | derivation | CLM-0924252480a1 | measured | explicit |  | 491abc0116d26daa | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-be77cedaa7a8
REL-4eac778b5096 | K143272 | CHN-K143272-SPO2 | CLM-73e2718fb82d | Oximetry sensor | component_of | CLM-54f39c86b89c | Pulse oximeter module: Nonin Xpod 3012 LP (low-power Xpod) | explicit |  | 96947be1f54673c3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-4eac778b5096
REL-e8538b92fd7a | K143272 | CHN-K143272-SPO2 | CLM-73e2718fb82d | Oximetry sensor | located_at | CLM-bf8541d0dc67 | Patient's finger | explicit |  | 2a57f5c1d3def738 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-e8538b92fd7a
REL-d937c4566a30 | K143272 | CHN-K143272-SPO2 | CLM-54f39c86b89c | Pulse oximeter module: Nonin Xpod 3012 LP (low-power Xpod) | located_at | CLM-e6276e6057cf | Xpod module fixed at the belt via clip | explicit |  | ed11a618ee001c41 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-d937c4566a30
REL-45a7027b1584 | K143272 | CHN-K143272-SPO2 | CLM-73e2718fb82d | Oximetry sensor | acquires | CLM-6d87e8e5044c | Not stated | not-stated | Raw optical signal not stated. | 08bf156da3b4559d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-45a7027b1584
REL-4b707019e851 | K143272 | CHN-K143272-SPO2 | CLM-73e2718fb82d | Oximetry sensor | measures | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | explicit | The document links oximeter to parameter directly, with no raw-signal description. | ba4d42abd199d721 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-4b707019e851
REL-f41bb8ced2d4 | K143272 | CHN-K143272-SPO2 | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | derivation | CLM-eca11ab54826 | measured | explicit |  | b83d23d7b6c2e7e2 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f41bb8ced2d4
REL-2b4f2c357fce | K143272 | CHN-K143272-PULSE | CLM-73e2718fb82d | Oximetry sensor | component_of | CLM-54f39c86b89c | Pulse oximeter module: Nonin Xpod 3012 LP (low-power Xpod) | explicit |  | 96947be1f54673c3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-2b4f2c357fce
REL-863efa9c35d4 | K143272 | CHN-K143272-PULSE | CLM-73e2718fb82d | Oximetry sensor | located_at | CLM-bf8541d0dc67 | Patient's finger | explicit |  | 2a57f5c1d3def738 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-863efa9c35d4
REL-c00259f98a30 | K143272 | CHN-K143272-PULSE | CLM-54f39c86b89c | Pulse oximeter module: Nonin Xpod 3012 LP (low-power Xpod) | located_at | CLM-e6276e6057cf | Xpod module fixed at the belt via clip | explicit |  | ed11a618ee001c41 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-c00259f98a30
REL-e9b6df898550 | K143272 | CHN-K143272-PULSE | CLM-73e2718fb82d | Oximetry sensor | acquires | CLM-6d87e8e5044c | Not stated | not-stated | Raw optical signal not stated. | 08bf156da3b4559d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-e9b6df898550
REL-9482b4f5381d | K143272 | CHN-K143272-PULSE | CLM-73e2718fb82d | Oximetry sensor | measures | CLM-92a1c7763280 | Pulse rate | explicit | The document links oximeter to parameter directly, with no raw-signal description. | ba4d42abd199d721 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-9482b4f5381d
REL-f473cb66c2df | K143272 | CHN-K143272-PULSE | CLM-92a1c7763280 | Pulse rate | derivation | CLM-eca11ab54826 | measured | explicit |  | b83d23d7b6c2e7e2 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f473cb66c2df
REL-77039e8cf531 | K143272 | CHN-K143272-SPO2 | CLM-e7ba308a04c1 | Average saturation | derived_feature_of | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | output-name | The output name refers to 'saturation'. | d1e9f2466490eb93 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-77039e8cf531
REL-4e9ae12bc163 | K143272 | CHN-K143272-SPO2 | CLM-5de192f8ed59 | Minimum saturation | derived_feature_of | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | output-name | The output name refers to 'saturation'. | 51f05881a0ef43d9 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-4e9ae12bc163
REL-a75e135ed3e4 | K143272 | CHN-K143272-SPO2 | CLM-83c2442cfd22 | Maximum saturation | derived_feature_of | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | output-name | The output name refers to 'saturation'. | c3fd9e523ea0bf79 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-a75e135ed3e4
REL-58f893b964a4 | K143272 | CHN-K143272-SPO2 | CLM-e1cd3ac8058c | Basal saturation | derived_feature_of | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | output-name | The output name refers to 'saturation'. | 6777ae9486a85f39 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-58f893b964a4
REL-f6524b41ed49 | K143272 | CHN-K143272-SPO2 | CLM-1cdb0340ac87 | ODI | reported_as | CLM-067ec2b7756b | Blood oxygen saturation (SpO2) | not-stated | 'ODI' is not expanded, and its input is not stated. | 0a4d1a667c728ddd | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f6524b41ed49
REL-7e9dd672ee4d | K143272 | CHN-K143272-PULSE | CLM-a9ce2d729da8 | Minimum Pulse Rate | derived_feature_of | CLM-92a1c7763280 | Pulse rate | output-name | The output name refers to 'Pulse Rate'. | e918896b530a3f25 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-7e9dd672ee4d
REL-96dfba31289c | K143272 | CHN-K143272-PULSE | CLM-5a17f0947641 | Maximum Pulse Rate | derived_feature_of | CLM-92a1c7763280 | Pulse rate | output-name | The output name refers to 'Pulse Rate'. | d341d856d9ac9f9a | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-96dfba31289c
REL-294db9928cd9 | K143272 | CHN-K143272-PULSE | CLM-f246d466ecda | Average Pulse Rate | derived_feature_of | CLM-92a1c7763280 | Pulse rate | output-name | The output name refers to 'Pulse Rate'. | 3ef2278ca446e66c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-294db9928cd9
REL-6cbb71532b6a | K143272 | CHN-K143272-POSITION | CLM-5db33eba2fd1 | Actigraphy sensor | measures | CLM-78792476b8f4 | Body position | explicit |  | e639da3e48ab0b62 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-6cbb71532b6a
REL-de9292cd2b97 | K143272 | CHN-K143272-POSITION | CLM-5db33eba2fd1 | Actigraphy sensor | located_at | CLM-32b036c02477 | Inside recorder worn on chest | inferred |  | 4ae16ebf1125d2fc | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-de9292cd2b97
REL-76acb392a787 | K143272 | CHN-K143272-POSITION | CLM-5db33eba2fd1 | Actigraphy sensor | acquires | CLM-78429ae1bb69 | Device tilt/rotation angle | inferred | The sensor is assumed to report the device’s angle; the source ties position to the device’s angle, not to the sensor. | 28b263ffa2179073 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-76acb392a787
REL-918954674a13 | K143272 | CHN-K143272-POSITION | CLM-78429ae1bb69 | Device tilt/rotation angle | represents | CLM-78792476b8f4 | Body position | explicit |  | 50d427453b7e718b | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-918954674a13
REL-15a346228962 | K143272 | CHN-K143272-POSITION | CLM-78792476b8f4 | Body position | derivation | CLM-a6f595eac9ad | Derived from device orientation (tilt/rotation angle) | explicit |  | 96e4ec5974e0608a | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-15a346228962
REL-0c2c4ddcaec2 | K143272 | CHN-K143272-POSITION | CLM-e0d8610a08dd | Not part of the ApneaLink Software report (explicit exclusion) | reported_as | CLM-78792476b8f4 | Body position | explicit | Negated: explicitly not reported. | 4369c337c2fd3d5c | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-0c2c4ddcaec2
REL-6cb3320c6ca7 | K143272 | CHN-K143272-SNORING | CLM-f923ac3ff49e | Not stated | measures | CLM-fae298439e0a | Snoring | not-stated | No snoring sensor described. | 7f1ee42ba9208c83 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-6cb3320c6ca7
REL-cdd1b63c4743 | K143272 | CHN-K143272-SNORING | CLM-f923ac3ff49e | Not stated | located_at | CLM-3b26fd752298 | Not stated | not-stated |  | 330685b63b76a067 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-cdd1b63c4743
REL-4194f2e1088a | K143272 | CHN-K143272-SNORING | CLM-f923ac3ff49e | Not stated | acquires | CLM-e00e5a54595a | Not stated | not-stated |  | 08bf156da3b4559d | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-4194f2e1088a
REL-9002a7ff7026 | K143272 | CHN-K143272-SNORING | CLM-fae298439e0a | Snoring | derivation | CLM-512244c68e62 | Not stated | not-stated |  | a7cd554f4410c077 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-9002a7ff7026
REL-e6d565d67631 | K143272 | CHN-K143272-SNORING | CLM-9bfd0bef8160 | Number of flow limited breaths with snoring | derived_feature_of | CLM-fae298439e0a | Snoring | output-name | The output name refers to snoring. | 023faecb1769e75b | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-e6d565d67631
REL-f31ae185de58 | K143272 | CHN-K143272-SNORING | CLM-7cdb877b6e99 | Number of flow limited breaths without snoring | derived_feature_of | CLM-fae298439e0a | Snoring | output-name | The output name refers to snoring. | b1985baa1b55c68e | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f31ae185de58
REL-58e053b05846 | K143272 | CHN-K143272-REPORT | CLM-964ea01d51e5 | AHI | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-58e053b05846
REL-f1eb52c1c5be | K143272 | CHN-K143272-REPORT | CLM-2102f8b48d70 | RI | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f1eb52c1c5be
REL-f3a040c70fe3 | K143272 | CHN-K143272-REPORT | CLM-f2dd379ff37c | Apnea Index | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-f3a040c70fe3
REL-7c5076655c06 | K143272 | CHN-K143272-REPORT | CLM-aeebd6206c8c | Hypopnea Index | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-7c5076655c06
REL-c8ec68c218b9 | K143272 | CHN-K143272-REPORT | CLM-ac3ad056bb9e | Cheyne-Stokes Breathing | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-c8ec68c218b9
REL-d6a59f0becdb | K143272 | CHN-K143272-REPORT | CLM-1cdb0340ac87 | ODI | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-d6a59f0becdb
REL-3b2dd7143a4e | K143272 | CHN-K143272-REPORT | CLM-7cdb877b6e99 | Number of flow limited breaths without snoring | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-3b2dd7143a4e
REL-a978d4250709 | K143272 | CHN-K143272-REPORT | CLM-9bfd0bef8160 | Number of flow limited breaths with snoring | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-a978d4250709
REL-0bf602dac074 | K143272 | CHN-K143272-REPORT | CLM-e7ba308a04c1 | Average saturation | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-0bf602dac074
REL-9d04dea91ba3 | K143272 | CHN-K143272-REPORT | CLM-5de192f8ed59 | Minimum saturation | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-9d04dea91ba3
REL-c43275a34821 | K143272 | CHN-K143272-REPORT | CLM-83c2442cfd22 | Maximum saturation | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-c43275a34821
REL-25689ec13345 | K143272 | CHN-K143272-REPORT | CLM-e1cd3ac8058c | Basal saturation | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-25689ec13345
REL-e1848ab60d7c | K143272 | CHN-K143272-REPORT | CLM-a9ce2d729da8 | Minimum Pulse Rate | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-e1848ab60d7c
REL-d3ac34cc4483 | K143272 | CHN-K143272-REPORT | CLM-5a17f0947641 | Maximum Pulse Rate | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-d3ac34cc4483
REL-27e1527fee54 | K143272 | CHN-K143272-REPORT | CLM-f246d466ecda | Average Pulse Rate | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 7c3486afdfa4e006 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-27e1527fee54
REL-991539a4af41 | K143272 | CHN-K143272-REPORT | CLM-78f52f4babfa | Apnea Index, unclassified | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 3a7ebf469a7627b3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-991539a4af41
REL-cfddfbdd05a4 | K143272 | CHN-K143272-REPORT | CLM-b0c786c21ea2 | Apnea Index, central | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 3a7ebf469a7627b3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-cfddfbdd05a4
REL-7b1506089132 | K143272 | CHN-K143272-REPORT | CLM-8cee7033615b | Apnea Index, mixed | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 3a7ebf469a7627b3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-7b1506089132
REL-38b9123b2cf8 | K143272 | CHN-K143272-REPORT | CLM-bcc4fe94e099 | Apnea Index, obstructive | generated_by | CLM-8ac34c8c8f77 | Report for the HCP generated by the ApneaLink Software | explicit |  | 3a7ebf469a7627b3 | pending | [6 empty] | https://vedantadc.github.io/evidence-explorer/510k/K143272/#item=REL-38b9123b2cf8

## Evidence anchors
Claim / relationship ID | Anchor ID | Role | Page(s) | Scope | Block / cell ID | Table row | Text layer | Exact quote
CLM-5f420010811a | ANC-e10754c594 | primary | 5 | document | K143272-V2-P005-B006 |  | native | The ApneaLink™ Air device is indicated for use by Health Care Professionals (HCP), where it may aid in the diagnosis of sleep disordered breathing for adult patients. ApneaLink Air records the following data: patient respiratory nasal airflow, snoring, blood oxygen saturation, pulse, respiratory effort and body position during sleep. The device uses these recordings to produce a report for the HCP that may aid in the diagnosis of sleep disordered breathing or for further clinical investigation. The device is intended for home and hospital use under the direction of a HCP.
CLM-5f420010811a | ANC-590da62197 | context | 3 | document | K143272-V2-P003-O004 |  | ocr_local_tesseract | ApneaLink Air records the following data: patient respiratory nasal airflow, snoring, blood oxygen saturation, pulse, respiratory effort and body position during sleep.
CLM-69b71b898d7c | ANC-a255a1baa3 | primary | 5 | document | K143272-V2-P005-B006 |  | native | where it may aid in the diagnosis of sleep disordered breathing for adult patients
CLM-39b492a99694 | ANC-20e8ef3e7c | context | 5 | document | K143272-V2-P005-B006 |  | native | where it may aid in the diagnosis of sleep disordered breathing for adult patients
CLM-006338992a5c | ANC-24e71a05dc | primary | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels
CLM-9feb9ae5aa69 | ANC-30f516f08e | primary | 5 | document | K143272-V2-P005-B003 |  | native | The disposable plastic nasal cannula is connected to the ApneaLink Air recorder and fixed at the patient’s nose.
CLM-fb528aac2fe3 | ANC-de3c112687 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | flow channel for measurement of patient respiratory nasal pressure.
CLM-fb528aac2fe3 | ANC-f7f381195f | context | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | that converts pressure changes to electrical signal levels
CLM-b7527e172867 | ANC-ea70c20a76 | primary | 5 | document | K143272-V2-P005-B006 |  | native | patient respiratory nasal airflow
CLM-b7527e172867 | ANC-a011db9681 | context | 10 | characteristic | K143272-V2-T02-R12-C1 | Location of linearization (Transformation of pressure values in flow values) | native | Location of linearization (Transformation of pressure values in flow values)
CLM-25cea7da6112 | ANC-818c0ce1d9 | limiting | 9 | predicate | K143272-V2-T02-R10-C2 | Sensor Technology | native | Respiratory Effort Sensors (RIP technology)
CLM-25cea7da6112 | ANC-b7346f702b | limiting | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
CLM-64a1213f973f | ANC-ce6a6871ef | primary | 5 | document | K143272-V2-P005-B003 |  | native | The ApneaLink Air recorder and the respiratory effort sensor must be fastened with the re-usable belt on the patient’s chest parallel to the spine.
CLM-64a1213f973f | ANC-cece337acf | context | 8 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | elastic cloth material belt to support unit and the respiratory effort sensor
CLM-e36feb851fad | ANC-a3694b07ca | context | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
CLM-cb110eefecb7 | ANC-11663e4746 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
CLM-cb110eefecb7 | ANC-c2349d64d5 | context | 5 | document | K143272-V2-P005-B006 |  | native | respiratory effort and body position during sleep
CLM-0924252480a1 | ANC-09431319de | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
CLM-73e2718fb82d | ANC-4e8329291b | primary | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes oximetry sensor
CLM-73e2718fb82d | ANC-e5fb61b3e8 | context | 5 | document | K143272-V2-P005-B003 |  | native | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger.
CLM-bf8541d0dc67 | ANC-548e31710a | primary | 5 | document | K143272-V2-P005-B003 |  | native | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger.
CLM-bf8541d0dc67 | ANC-93989da872 | context | 8, 9 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | Flexi Wrap to fix the oximeter sensor to the patient’s finger.
CLM-b02ad0500a1b | ANC-e28d1d9dd0 | context | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
CLM-f706d25f6cda | ANC-a540d69e21 | context | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
CLM-eca11ab54826 | ANC-6014eaa460 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
CLM-5db33eba2fd1 | ANC-d6885708f1 | primary | 9, 10 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes solid-state actigraphy sensor for position
CLM-5db33eba2fd1 | ANC-bcce02dd46 | context | 5 | document | K143272-V2-P005-B010 |  | native | Actigraphy sensor provides body position of patient during sleep.
CLM-32b036c02477 | ANC-daf92dd798 | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
CLM-32b036c02477 | ANC-86db247940 | context | 13 | subject | K143272-V2-T02-R18-C4 | IEC 60601-1 (Electrical safety) | native | The actigraphy sensor was already physically incorporated in ApneaLink Pro.
CLM-32b036c02477 | ANC-e54c097820 | context | 5 | document | K143272-V2-P005-B003 |  | native | The ApneaLink Air recorder and the respiratory effort sensor must be fastened with the re-usable belt on the patient’s chest parallel to the spine.
CLM-a5dcc1a3a694 | ANC-7a75ac86cc | context | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
CLM-78792476b8f4 | ANC-185c1dd57a | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | body position channel to determine body position of the patient during sleep
CLM-964ea01d51e5 | ANC-9a7b444af0 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | AHI
CLM-964ea01d51e5 | ANC-336db180fc | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-2102f8b48d70 | ANC-f6bd40871c | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | RI
CLM-2102f8b48d70 | ANC-156eaad8e3 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-f2dd379ff37c | ANC-a6c8419b99 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
CLM-f2dd379ff37c | ANC-1ddbae5993 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-aeebd6206c8c | ANC-92c0ddb761 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Hypopnea Index
CLM-aeebd6206c8c | ANC-6c51ab4019 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-7cdb877b6e99 | ANC-fb1a449e16 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Number of flow limited breaths without snoring
CLM-7cdb877b6e99 | ANC-f8d48ddd4c | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-9bfd0bef8160 | ANC-0988f6fb9a | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Number of flow limited breaths with snoring
CLM-9bfd0bef8160 | ANC-7609011750 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-ac3ad056bb9e | ANC-d787e379f6 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Cheyne-Stokes Breathing
CLM-ac3ad056bb9e | ANC-2dc5362570 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-1cdb0340ac87 | ANC-13dc1d193a | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | ODI
CLM-1cdb0340ac87 | ANC-4da95ab8fe | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-e7ba308a04c1 | ANC-b277704478 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Average saturation
CLM-e7ba308a04c1 | ANC-e7aad13982 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-5de192f8ed59 | ANC-f20bb19861 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Minimum saturation
CLM-5de192f8ed59 | ANC-e3c6911ad7 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-83c2442cfd22 | ANC-a39b1e70f9 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Maximum saturation
CLM-83c2442cfd22 | ANC-6b2d0185ac | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-e1cd3ac8058c | ANC-fb82983abe | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Basal saturation
CLM-e1cd3ac8058c | ANC-3b39714093 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-a9ce2d729da8 | ANC-680754ef4e | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Minimum Pulse Rate
CLM-a9ce2d729da8 | ANC-2472638c88 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-5a17f0947641 | ANC-d3840b9f8c | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Maximum Pulse Rate
CLM-5a17f0947641 | ANC-a65a310fc1 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-f246d466ecda | ANC-de71ecbd61 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Average Pulse Rate
CLM-f246d466ecda | ANC-1ac87b3f95 | context | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
CLM-030e4da72b91 | ANC-a3f0d7069c | limiting | 5 | document | K143272-V2-P005-B010 |  | native | Data is neither part of the report generated by ApneaLink Software nor does it influence or alter any index of the report.
CLM-77cb9c564d4e | ANC-ddafdfd8ef | primary | 8 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | Plastic tubing and cannula for pressure sensing
CLM-77cb9c564d4e | ANC-baeea33bc1 | context | 5 | document | K143272-V2-P005-B003 |  | native | The disposable plastic nasal cannula is connected to the ApneaLink Air recorder and fixed at the patient’s nose.
CLM-a7cf317deee5 | ANC-58d409087a | context | 5 | document | K143272-V2-P005-B003 |  | native | The disposable plastic nasal cannula is connected to the ApneaLink Air recorder and fixed at the patient’s nose.
CLM-28294e7255aa | ANC-b03c7c80c7 | primary | 10 | subject | K143272-V2-T02-R12-C4 | Location of linearization (Transformation of pressure values in flow values) | native | ApneaLink Air recorder
CLM-28294e7255aa | ANC-cefe789e6b | primary | 10 | characteristic | K143272-V2-T02-R12-C1 | Location of linearization (Transformation of pressure values in flow values) | native | Location of linearization (Transformation of pressure values in flow values)
CLM-0be59426dade | ANC-73ee841b54 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
CLM-0be59426dade | ANC-cc703c9d3a | context | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels
CLM-0be59426dade | ANC-c3d2a3760b | context | 5 | document | K143272-V2-P005-B003 |  | native | The respiratory effort sensor is connected to the ApneaLink Air recorder and held in place by the belt.
CLM-5f8f25c65423 | ANC-db28adbe9c | primary | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels
CLM-5f8f25c65423 | ANC-f8a6e99ed8 | context | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
CLM-54f39c86b89c | ANC-fe2408cbcd | primary | 13 | subject | K143272-V2-T02-R21-C4 | Xpod oximeter | native | Nonin Xpod 3012 LP
CLM-54f39c86b89c | ANC-bb034aeb42 | context | 5 | document | K143272-V2-P005-B012 |  | native | Low power Xpod replaces previous Xpod model.
CLM-e6276e6057cf | ANC-f850b3b027 | primary | 8 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | The new low power Xpod is fixed at the belt via clip.
CLM-6d87e8e5044c | ANC-c59aa32381 | context | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
CLM-067ec2b7756b | ANC-0f30efa870 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
CLM-067ec2b7756b | ANC-845e8d36d6 | context | 5 | document | K143272-V2-P005-B006 |  | native | blood oxygen saturation
CLM-92a1c7763280 | ANC-b0e70715e6 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
CLM-92a1c7763280 | ANC-5f99f35e81 | context | 5 | document | K143272-V2-P005-B006 |  | native | blood oxygen saturation, pulse, respiratory effort
CLM-a8c39fbe3950 | ANC-6d5bb9f867 | primary | 5 | document | K143272-V2-P005-B012 |  | native | Same accuracy and performance as the existing Xpod, but considerably lower power draw.
CLM-a8c39fbe3950 | ANC-d92bdd196d | context | 14 | comment | K143272-V2-T02-R21-C5 | Xpod oximeter | native | Xpod 3012 and Xpod 3012 LP have equivalent performance and clinical accuracy.
CLM-78429ae1bb69 | ANC-418c2bf5db | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
CLM-a6f595eac9ad | ANC-72b6d97549 | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
CLM-e0d8610a08dd | ANC-4ffa370a17 | primary | 5 | document | K143272-V2-P005-B010 |  | native | Data is neither part of the report generated by ApneaLink Software nor does it influence or alter any index of the report.
CLM-e0d8610a08dd | ANC-fc6f05652a | context | 11 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Recorded body position data can be viewed using an EDF viewer but is not displayed in the report generated by the ApneaLink Software.
CLM-7ec48c11db89 | ANC-b2d7372c08 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The body position data is stored as an EDF+ file on the recorder. It is accessible on a computer via USB connection using an EDF viewer.
CLM-55d0a82ea239 | ANC-5073d01ed9 | primary | 15 | document | K143272-V2-P015-B010 |  | native | Additional side by side testing has been performed by using a protractor to show the defined position around the body axis for ApneaLink Air and Nox T3.
CLM-fae298439e0a | ANC-97a78ea7a1 | primary | 5 | document | K143272-V2-P005-B006 |  | native | patient respiratory nasal airflow, snoring, blood oxygen saturation
CLM-f15b1f62480b | ANC-0161181ebf | primary | 10 | subject | K143272-V2-T02-R13-C4 | Interface between embedded software and PC software/EDF viewer | native | Recorded data is stored in European Data Format (EDF+) file on SD card. When device is connected to PC via USB the device provides access to its internal memory as mass storage memory including the recorded EDF+ files.
CLM-f15b1f62480b | ANC-1cbddb7483 | context | 10, 11 | subject | K143272-V2-T02-R13-C4 | Interface between embedded software and PC software/EDF viewer | native | Data related to patient respiratory nasal airflow, snoring, blood oxygen saturation, pulse, respiratory effort can be transferred to the ApneaLink Software.
CLM-0e8dbbca0f00 | ANC-5c74323d1b | primary | 15 | document | K143272-V2-P015-B002 |  | native | Detection of respiratory events and reported indices was cleared under K083575 (ApneaLink Plus) and remained unchanged in ApneaLink Pro (K131932) and ApneaLink Air.
CLM-0e8dbbca0f00 | ANC-c61e2eab9d | context | 15 | document | K143272-V2-P015-B002 |  | native | Thus side-by-side testing results from ApneaLink Pro comparing detection of all respiratory events remain valid.
CLM-8ac34c8c8f77 | ANC-e8af800caf | primary | 5 | document | K143272-V2-P005-B006 |  | native | The device uses these recordings to produce a report for the HCP
CLM-8ac34c8c8f77 | ANC-89881e4b7a | context | 5 | document | K143272-V2-P005-B003 |  | native | With the ApneaLink Software the physician can generate a report with the recorded and analyzed data (respiratory pressure, respiratory effort, pulse rate, oxygen saturation) to aid in diagnosis.
CLM-78f52f4babfa | ANC-ef2d3dbbad | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
CLM-b0c786c21ea2 | ANC-6618316043 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
CLM-8cee7033615b | ANC-9ba6dee649 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
CLM-bcc4fe94e099 | ANC-dfd4e25fa0 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
REL-99a63111d1a2 | ANC-2c8fdabfa5 | primary | 8 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | Plastic tubing and cannula for pressure sensing
REL-6669ab69f70e | ANC-1e2737d4f9 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The disposable plastic nasal cannula is connected to the ApneaLink Air recorder and fixed at the patient’s nose.
REL-4612222a5926 | ANC-d41187243d | primary | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels
REL-4612222a5926 | ANC-2329c38235 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | flow channel for measurement of patient respiratory nasal pressure.
REL-7fa8540c4751 | ANC-6a081fd88e | primary | 10 | characteristic | K143272-V2-T02-R12-C1 | Location of linearization (Transformation of pressure values in flow values) | native | Location of linearization (Transformation of pressure values in flow values)
REL-7fa8540c4751 | ANC-35252fa70d | primary | 10 | subject | K143272-V2-T02-R12-C4 | Location of linearization (Transformation of pressure values in flow values) | native | ApneaLink Air recorder
REL-9734ff192882 | ANC-054f1881a3 | primary | 10 | subject | K143272-V2-T02-R12-C4 | Location of linearization (Transformation of pressure values in flow values) | native | ApneaLink Air recorder
REL-79d1e425d13a | ANC-f2f21cf7dd | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Number of flow limited breaths without snoring
REL-e8498d05cc18 | ANC-1205bfec27 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Number of flow limited breaths with snoring
REL-e8750cd5dc13 | ANC-2e3821fb03 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The ApneaLink Air recorder and the respiratory effort sensor must be fastened with the re-usable belt on the patient’s chest parallel to the spine.
REL-70a59a69c068 | ANC-53d82da520 | primary | 9 | subject | K143272-V2-T02-R10-C4 | Sensor Technology | native | Utilizes solid-state pressure sensor for nasal flow and effort that converts pressure changes to electrical signal levels
REL-b58c7986c74e | ANC-476095d6f9 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
REL-be77cedaa7a8 | ANC-ba9f4e11b2 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | effort channel to measure the respiratory effort using a pneumatic principle.
REL-4eac778b5096 | ANC-e9c612b7d5 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger.
REL-e8538b92fd7a | ANC-42b5db3919 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger.
REL-d937c4566a30 | ANC-14c6f304a0 | primary | 8 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | The new low power Xpod is fixed at the belt via clip.
REL-4b707019e851 | ANC-df93725d01 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
REL-f41bb8ced2d4 | ANC-cf55f0078c | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
REL-2b4f2c357fce | ANC-74305c2c09 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger.
REL-863efa9c35d4 | ANC-bec69d1a29 | primary | 5 | document | K143272-V2-P005-B003 |  | native | The oximetry sensor is connected to the Xpod and fixed at the patient’s finger.
REL-c00259f98a30 | ANC-4bb36a62dd | primary | 8 | subject | K143272-V2-T02-R05-C4 | Method of connection to the Patient | native | The new low power Xpod is fixed at the belt via clip.
REL-9482b4f5381d | ANC-e4b60605c9 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
REL-f473cb66c2df | ANC-12a6b3a686 | primary | 8 | subject | K143272-V2-T02-R04-C4 | Channels | native | oximeter input to measure degree of oxygen saturation of the blood and pulse rate.
REL-77039e8cf531 | ANC-edea213688 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Average saturation
REL-4e9ae12bc163 | ANC-4b82f65921 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Minimum saturation
REL-a75e135ed3e4 | ANC-7d4ebb6423 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Maximum saturation
REL-58f893b964a4 | ANC-f0a6e0ba0c | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Basal saturation
REL-7e9dd672ee4d | ANC-14cc5b761b | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Minimum Pulse Rate
REL-96dfba31289c | ANC-d7b9127e84 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Maximum Pulse Rate
REL-294db9928cd9 | ANC-b899211639 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Average Pulse Rate
REL-6cbb71532b6a | ANC-9b6dcb3b87 | primary | 5 | document | K143272-V2-P005-B010 |  | native | Actigraphy sensor provides body position of patient during sleep.
REL-de9292cd2b97 | ANC-541712ce8c | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
REL-de9292cd2b97 | ANC-122ae588b4 | primary | 13 | subject | K143272-V2-T02-R18-C4 | IEC 60601-1 (Electrical safety) | native | The actigraphy sensor was already physically incorporated in ApneaLink Pro.
REL-76acb392a787 | ANC-6edba0aa9f | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
REL-918954674a13 | ANC-fb7f74416e | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
REL-15a346228962 | ANC-2d15128c55 | primary | 8 | comment | K143272-V2-T02-R04-C5 | Channels | native | The ApneaLink Air device shows similar to the Nox T3 body position information according to the angle the device is tilted or rotated.
REL-0c2c4ddcaec2 | ANC-58702f387e | primary | 5 | document | K143272-V2-P005-B010 |  | native | Data is neither part of the report generated by ApneaLink Software nor does it influence or alter any index of the report.
REL-e6d565d67631 | ANC-8c4f8e4e96 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Number of flow limited breaths with snoring
REL-f31ae185de58 | ANC-012f05224d | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Number of flow limited breaths without snoring
REL-58e053b05846 | ANC-0f7a41816b | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-58e053b05846 | ANC-f58e30a9fa | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-f1eb52c1c5be | ANC-e4c730de7e | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-f1eb52c1c5be | ANC-bb79fbe941 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-f3a040c70fe3 | ANC-ffce349f16 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-f3a040c70fe3 | ANC-d7d3b88c5a | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-7c5076655c06 | ANC-d09cbb90e3 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-7c5076655c06 | ANC-1b2f687a92 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-c8ec68c218b9 | ANC-e15bb84be5 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-c8ec68c218b9 | ANC-0c251ce69a | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-d6a59f0becdb | ANC-f5b1912d0b | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-d6a59f0becdb | ANC-43750437bb | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-3b2dd7143a4e | ANC-cc9b88e219 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-3b2dd7143a4e | ANC-6c757d0afa | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-a978d4250709 | ANC-12b2096ab5 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-a978d4250709 | ANC-78233ac8c1 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-0bf602dac074 | ANC-4bd51e3dfc | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-0bf602dac074 | ANC-4fba9a273b | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-9d04dea91ba3 | ANC-2b9b98c316 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-9d04dea91ba3 | ANC-333528fbd6 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-c43275a34821 | ANC-07c3046716 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-c43275a34821 | ANC-28268fd787 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-25689ec13345 | ANC-8033358070 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-25689ec13345 | ANC-27c81b96b3 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-e1848ab60d7c | ANC-d2fceddc7d | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-e1848ab60d7c | ANC-06826d9b5f | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-d3ac34cc4483 | ANC-e6a22206b8 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-d3ac34cc4483 | ANC-523bfbad7c | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-27e1527fee54 | ANC-4107cb2163 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | The following indices are generated from the ApneaLink Software:
REL-27e1527fee54 | ANC-4ba2ebe5a0 | primary | 11, 12 | comment | K143272-V2-T02-R15-C5 | Analysis result (Indices) | native | Thus the indices generated from the ApneaLink Software in the report remain unchanged.
REL-991539a4af41 | ANC-b7d788ded5 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
REL-cfddfbdd05a4 | ANC-6afa448715 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
REL-7b1506089132 | ANC-30d36cd9cb | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)
REL-38b9123b2cf8 | ANC-1514621cb1 | primary | 11 | shared | K143272-V2-T02-R15-C3-C4 | Analysis result (Indices) | native | Apnea Index (unclassified, central, mixed, obstructive)

## Table 8.3 (verified)
Physiological parameter | Parameter N | Sensor type | Sensor-row N | Sensor location(s) | Raw/acquired signal | Direct / derived | Derived/reported features | Explicit exclusions
No verified rows. Table 8.3 is empty until claims and relationships are human-verified.